File Revision Date:

---- -------- -----

July 24, 2026

Data Set Description:

---- --- ------------

PI: 		Ryan M. Stauffer (prior to May 2021, PI: Anne M. Thompson)

Instrument: 	ECC Ozonesondes

Site: 		Natal, Brazil (Brazilian Space Agency - INPE, NASA Goddard and Wallops Flight Facility)

Latitude: 	5.84S

Longitude: 	35.21W

Altitude: 	42m amsl

Measurement Quantities: Ozone partial pressure, Ozone mixing ratio, Pressure, Temperature,

Relative humidity, Geopotential height, GPS Altitude, Latitude and Longitude of payload,

and Wind.

Data Version Description: SHADOZ V06 reprocessed/homogenized data.

Archived Data Record (SHADOZ V06) Start Date: August 23, 2000

Details of V06 reprocessing are in the references below.

Contact Information:

------- ------------

Name:		Dr. Ryan M. Stauffer

Address:	NASA Goddard Space Flight Center (GSFC)

8800 Greenbelt Road, Mail Code 614

Greenbelt, Maryland 20771 U.S.A.

Phone:		1(301)614-5552

Email:  		ryan.m.stauffer@nasa.gov

Name:		Dr. Anne M. Thompson

Address:	NASA Goddard Space Flight Center (GSFC)

8800 Greenbelt Road, Mail Code 614

Greenbelt, Maryland 20771 U.S.A.

Email:  		anne.m.thompson@nasa.gov

Site Contact:

----- -------

Name:		Dr. Francisco Raimundo da Silva

Address:	Environmental Variables Laboratory

Brazilian National Institute of Space Research

Instituto Nacional de Pesquisas Espaciais (INPE)

Natal, Brazil

Email:		francisco.raimundo@inpe.br

Data Contacts:

---- --------

SHADOZ GSFC Processing:

Name:		Dr. Debra E. Kollonige

Address:	NASA Goddard Space Flight Center (Code 614)

8800 Greenbelt Road

Greenbelt, Maryland 20771 U.S.A.

Email:		debra.e.kollonige@nasa.gov

WFF Processing/QA (record prior to Feb 2025):

Name:		E. Thomas Northam

Address:	NASA Goddard Space Flight Center / Wallops Flight Facility (Code 612)

Wallops Island, Virginia 23337 U.S.A.

Email:		e.thomas.northam@nasa.gov

DOI:

----

https://doi.org/10.57721/shadoz-v06

https://doi.org/10.57721/shadoz-v06-natal

Data License:

---- --------

CC0

Data L0/L1 Location:

-------------

NASA GSFC Code 614 Cluster

Reference Articles:

--------- ---------

Witte, J.C., A. M. Thompson, H. G. J. Smit, M. Fujiwara, F. Posny, Gert J. R. Coetzee, ... F. R. da Silva (2017), First reprocessing of Southern Hemisphere ADditional OZonesondes (SHADOZ) profile records (1998-2015): 1. Methodology and evaluation, J. Geophys. Res. Atmos., 122, 6611-6636. https://doi.org/10.1002/2016JD026403.

Thompson, A. M., J. C. Witte, C., Sterling, A., Jordan, B. J., Johnson, S. J. Oltmans, ... Thiongo, K. (2017). First reprocessing of Southern Hemisphere Additional Ozonesondes (SHADOZ) ozone profiles (1998-2016): 2. Comparisons with satellites and ground-based instruments. Journal of Geophysical Research: Atmospheres, 122, 13,000-13,025. https://doi.org/10.1002/2017JD027406.

Witte, J. C., Thompson, A. M., Smit, H. G. J., Vomel, H., Posny, F., & Stuebi, R. (2018). First reprocessing of Southern Hemisphere ADditional OZonesondes profile records: 3. Uncertainty in ozone profile and total column. Journal of Geophysical Research: Atmospheres, 123, 3243-3268. https://doi.org/10.1002/2017JD027791.

Instrument Description:

---------- ------------

The ECC Ozonesonde (Electrochemical Concentration Cell Ozonesonde) is a lightweight,

balloon-borne instrument mated to a meteorological radiosonde and flown to 30+ km while

transmitting data back to a ground station. The heart of the ozonesonde is an electochemical

concentration cell (ECC) that senses ozone as it reacts with a dilute solution of potassium iodide

to produce an electrical current proportional to the ozone concentration of the air.

Project start date: 		August 1979

Start digital data acquisition: 	July 24, 1996

Data record: 			08/2000-current

Data gaps: 	05/2011-09/2013, 06/2022-09/2022; COVID-19 shutdowns: 04/2020-10/2020, 04/2021, 07/2021, 01/2022–02/2022, 04/2022; 03/2025-06/2025, 04/2026-06/2026

Ozone sensor:

Science Pump Corporation (SPC) ECC-6A (Entire record)

Environmental Science (EN-SCI) Corporation 1 or 2Z (99 records 2000-2002)

Radiosonde:

Sippican, Inc. ZEEMET Mark II MICROSONDE 403 MHz dGPS with Ozone (2000–12/2003)

Sippican, Inc. ZEEMET Mark IIA MICROSONDE 403 MHz dGPS with Ozone (12/2003–03/2007)

Lockheed Martin Mark IIA MICROSONDE 403 MHZ DGPS (LOS) with Ozone Sensor (10/2006–08/2016)

Lockheed Martin: LMS6 Differential GPS (LOS) 403 MHz (08/2016–02/2025)

Vaisala: RS41-SG (02/2025 – present)

Sensing Solution Type (SST):

1% KI, 1.0x (full) buffer (Entire record w/ SPC)

1% KI, 1.0x (full) buffer (99 records w/ EN-SCI, transfer function used to convert data to SPC 1%, full buffer)

Launch Frequency:

------ ----------

Local Launch times: nominal 1300-1500 UTC weekly, since COVID-19 shutdowns 2-3 times per month frequency

Algorithm Description:

--------- ------------

Ozone is calculated as a partial pressure. PTU data from the sonde

is not used directly in the calculation except in the pump correction.

PPOZ(nb) = 0.004307 * i * Temperature * t *pcf

where:

the constant is half the ratio of ideal gas constant to Faraday's constant.

i is the current from the sensor - background in uA.

t is the time in seconds to pump 100 CCs of air through the pump.

Temperature is the pump temperature (K).

Pcf is the pump correction factor to account for loss in pump efficiency at lower

pressures.

Pre-flight procedures comply with:

"WMO/GAW Report 201"

https://library.wmo.int/doc_num.php?explnum_id=7167

"WMO/GAW Report 268"

https://library.wmo.int/index.php?lvl=notice_display&id=21986#.YaFNSbpOlc8

All data have been reprocessed to comply with:

"O3S-DQA-Guidelines Homogenization-V2-19November2012.pdf"

http://www-das.uwyo.edu/~deshler/NDACC_O3Sondes/O3s_DQA/O3S-DQA-Guidelines%20Homogenization-V2-19November2012.pdf

Expected Precision/Accuracy of Instrument:

-------- ------------------ -- -----------

Ozonesonde:

Accuracy	Precision	Resolution

+/- 5%		+/- 4%		~150m

Instrument History:

---------- --------

ECC Ozonesondes:

Manufacturer:	   Model # 	Time Period Used	Design and Changes

Science Pump       6A	2000-present		Pump temp. thermistor inside of pump block

EN-SCI		   Z		 05/2000-07/2002	Different manufacturer but same design as 6A

Solution Recipe changes:

2000-present -> 1% KI, 1.0x (full) buffer (Entire record w/ SPC)

2000-2002     ->	1% KI, 1.0x (full) buffer (99 records w/ EN-SCI, transfer function used to convert data to SPC 1%, full buffer)

Radiosonde changes:

2000		-> Sippican, Inc. ZEEMET Mark II MICROSONDE 403 MHz dGPS with Ozone sensor

12/2003	-> Sippican, Inc. ZEEMET Mark IIA MICROSONDE 403 MHz dGPS with Ozone sensor

10/2006	-> Lockheed Martin Mark IIA MICROSONDE 403 MHZ DGPS (LOS) with Ozone Sensor

08/2016	-> Lockheed Martin LMS6 Differential GPS (LOS) 403 MHz with Ozone sensor

02/2025 	-> Vaisala: RS41-SG (started officially in February with a few test flights starting in 2024)

Solution volume changes:

2000-present -> 3.0 cm^3 used, no correction needed